## Slide 1: GEM5000 Stress Test Report
Haining Bao

## Slide 2: Stress Test Configuration
20 GEM5000’s:
  2GB RAM
  Desktop
  G5K v1.2.0 B148
Network Up/Down Scripts:
  4 each @ 5, 5, 10, 40min
  2 @ random intervals
Sampling rate
  29/hr < Rate < 30/hr
Run Time: 16 Days
GWP Server
  G8 Server
  4 CPUs (2GHz each)
  4GB RAM
  Virtual Machine
  GWP 5.1 upgraded from B36 to B40(Release Build) on 6/9/2017

## Slide 3: Failure Trend G5K 1.3.0
Chart: 
| GWP B36 G5K B149 | GWP B40 G5K B1 | GWP B40 G5K B2 | GWP B40 G5K B5 | GWP B40 G5K B9
Stable | 0 | 0.1 | 0.1429 | 0 | 0.0474
Unstable | 0.0222 | 0.0556 | 0.0238 | 0 | 0.0104
Overall | 0.0067 | 0.0167 | 0.1053 | 0 | 0.035

## Slide 4: Failure Trend G5K 1.2.0
Chart: 
| GWP B18 G5K B96 | GWP B26 G5K B119 | GWP B26 G5K B143 | GWP B30 G5K B146 | GWP B30 G5K B148 | GWP B36 G5K B149
Stable | 0.42 | 0.1 | 0.1212 | 0.0268 | 0 | 0
Unstable | 0.47 | 0.115 | 0.1228 | 0.0476 | 0 | 0.0222
Overall | 0.44 | 0.1075 | 0.1217 | 0.0325 | 0 | 0.0067

## Slide 5: Failures and Frequency
 | Tickets |  |  |  |  | 
CR Number | IL96-7445 | IL96-7446/7400 | IL96-7394 | IL96-7138 | IL96-6851 | IL96-7074
Stable Network | 3 | 4 | 2 | 7 | 4 | 0
Unstable Network | 0 | 0 | 0 | 2 | 1 | 1
Overall | 2 | 4 | 2 | 9 | 0 | 0
IL96-7445 – Mismatch sample# between GWP and Clients
IL96-7446/7400 – Wrong network icon color
IL96-7394 – UI freeze
IL96-7138 – Out of Memory issue
IL96-6851 – 3206 crash similar but different flavor of JMS issue
IL96-7074 – AM eeprom error handling related issue
Notes: Both CRs have been resolved and are in the process of being tested