## Slide 1: GEM 4000 Stress Test Configuration - Current
Product | SBC | H.W. | Ver. | Stable | Unstable | Total #
G4K | New Seco | Product | V3.1 b24 | 1 | 10 | 11
G4K | Old Seco | Product | V3.1 b24 | 1 | 17 | 18
G4K | Kontron | Product | V3.1 b24 | 1 | 2 | 3
Server | G8 | Virtual | V5.1.0 b38 | 1 | N/A | 1
Note: Sample frequency is 16 samples an hour.

## Slide 2: GEM 4000 Failure Rate Trend
GWP5.1.0 b36, G4K b24 | Stable | Unstable | Overall
Baseline | 0.0% | 0.4% | 0.3%
Chart: 
| G4k, release; GWP, b40 (7 Days) | G4k, release; GWP, b38 (5 Days) | G4k, release; GWP, b36 (8 Days)
Stable | 0.0417 | 0.375 | 0
Unstable | 0.0733 | 0.1173 | 0.0339
Overall | 0.0703 | 0.1404 | 0.0299

## Slide 3: GEM 4000 Failures and Frequency – b36
CR# | Type | Summary | Frequency
18622 | Crash | Multiple 3206 on analyzer before restart. 1:00 AM. | 5
GWP 510
Error Type | New | Kontron | Old
Crash | 0 | 0 | 0
Freeze | 0 | 0 | 0
Stable Network
Error Type | New | Kontron | Old
Crash | 1 | 1 | 4
Freeze | 0 | 0 | 1
Unstable Network

## Slide 4: GEM 4000 Failures and Frequency – b38
CR# | Type | Summary | Frequency
18622 | Crash | Multiple 3206 on analyzer before restart. 1:00 AM. | 25
GWP 510
Error Type | New | Kontron | Old
Crash | 3 | 2 | 1
Freeze | 0 | 0 | 0
Stable Network
Error Type | New | Kontron | Old
Crash | 12 | 0 | 7
Freeze | 0 | 0 | 0
Unstable Network

## Slide 5: GEM 4000 Failures and Frequency – b40
CR# | Type | Summary | Frequency
18622 | Crash | Multiple 3206 on analyzer before restart. 1:00 AM. | 13
GWP 510
Error Type | New | Kontron | Old
Crash | 0 | 1 | 0
Freeze | 0 | 0 | 0
Stable Network
Error Type | New | Kontron | Old
Crash | 7 | 3 | 2
Freeze | 1 | 0 | 4
Unstable Network