## Slide 1: GEM5000 Stress Test Report
Haining Bao

## Slide 2: Stress Test Configuration
20 GEM5000’s:
  2GB RAM
  Desktop
  G5K v1.3.0 B29
Network Up/Down Scripts:
  4 each @ 5, 5, 10, 40min
  2 @ random intervals
Sampling rate
  29/hr < Rate < 30/hr
Run Time: 27 Days
GWP Server
  G8 Server
  4 CPUs (2GHz each)
  4GB RAM
  Virtual Machine
  Upgraded from GWP 5.1 B40(Release Build) to 5.2 B9 on 8/11/2017

## Slide 3: Failure Trend G5K 1.3.0
Chart: 
| GWP-B36 G5K-B149 | GWP-B40 G5K-B1 | GWP-B40 G5K-B2 | GWP-B40 G5K-B5 | GWP-B40 G5K-B9 | GWP-B40 G5K-B23 | GWP-5.2-B9 G5K-B29
Stable | 0 | 0.1 | 0.1429 | 0 | 0.0474 | 0.0786 | 0
Unstable | 0.0222 | 0.0556 | 0.0238 | 0 | 0.0104 | 0 | 0.0062
Overall | 0.0067 | 0.0167 | 0.1053 | 0 | 0.035 | 0.0529 | 0.0019

## Slide 4: Failures and Frequency
 | Ticket
CR Number | IL96-7394
Stable Network | 0
Unstable Network | 1
Overall | 1
Issues under monitoring:
IL96-7394 – UI freeze
IL96-7226 – Client Copy IL Data time out issue
IL96-7074 – AM eeprom error handling related issue
Notes: Both CRs have been resolved and are in the process of being tested